Step Noise Finder

User Guide

Version 1.2.0

Automated Step Noise Detection for TDD Baseline Files

Nabsys HD-Mapping TDD File Analysis

Python / Tkinter / Matplotlib / SciPy

# 1. Introduction

The Step Noise Finder is a standalone desktop application that automatically scans every channel of one or more TDD baseline files and identifies regions containing step noise. It is designed for Nabsys HD-Mapping instruments that produce multi-channel voltage data (typically 256 channels).

Step noise is a characteristic interference pattern where the signal abruptly rises, holds at an elevated level, then falls back, repeating in a periodic pattern. The application detects step noise in the 4–20 Hz frequency range with configurable amplitude thresholds, and offers a separate envelope mode for catching non-periodic bursts of activity in the same band.

## 1.1 What's New in 1.2.0

- Envelope-mode detection — catches bursty events that aren't strictly periodic.

- Cross-file merging — a step-noise event that spans several consecutive TDD files is reported as a single event.

- Recursive Scan All Runs — point at a top-level data root and the app walks the tree, finding every folder containing *Baseline*.tdd files.

- Channel-level parallelism — scans use all CPU cores; combined with FFT autocorrelation this is roughly 10–30× faster than the previous serial implementation.

- Sortable result columns and a per-folder Sherlock-compatible JSON report alongside the existing text report.

- Saved settings — the toolbar parameters persist between runs in step_noise_finder_config.json.

- 8 GB bounded LRU cache — large batch scans no longer crash from memory pressure; files are reloaded transparently if you click a result whose data was evicted.

# 2. Getting Started

## 2.1 Running the Application

From a checked-out source tree using the project venv:

.venv\Scripts\python.exe src\step_finder.py

Or use the standalone executable build:

StepNoiseFinder.exe

## 2.2 System Requirements

- Python 3.13 with NumPy, SciPy, and Matplotlib (or the standalone EXE — no Python required)

- Optional: the sherlock package (Nabsys-Organization/sherlock on GitHub) for JSON RunResults output. Without it the .json sibling reports are silently skipped; the .txt reports are always produced.

- Windows 10/11

- 8 GB RAM recommended for large batch scans (the app self-limits its file cache to 8 GB)

# 3. User Interface

The window has four areas: a menu bar, a parameter toolbar, a results table on the left, and a plot panel on the right. A status line at the bottom shows progress and report-write notifications.

## 3.1 File Menu

| Menu Item | Shortcut | Description |
| --- | --- | --- |
| Open TDD… | Ctrl+O | Open one TDD file. Press Scan All Channels afterward to analyze it. |
| Scan Baselines… | Ctrl+Shift+O | Custom folder picker showing the *Baseline*.tdd files in the chosen folder. Scans every Baseline TDD in that folder. |
| Scan Files… |  | Pick individual TDD files (any *.tdd) and scan only those. |
| Scan Folders… |  | Add multiple folders to a queue (one at a time, with previews) and scan them in batch. Each folder produces its own report. |
| Scan All Runs… |  | Pick a single root directory; the app recursively finds every subfolder containing *Baseline*.tdd files. A preview dialog lets you remove rows before scanning. |
| Exit |  | Close the application. Settings are saved on every parameter change, so nothing is lost. |

## 3.2 Toolbar

All toolbar values persist between sessions (saved to step_noise_finder_config.json next to the app) and are applied when you click Scan All Channels or any of the menu commands.

| Control | Default | Description |
| --- | --- | --- |
| File | — | Loaded file or batch label (filename truncated before the first underscore for readability). |
| Amp ≥ (mV) | 100 | Minimum peak-to-peak amplitude of the bandpass signal in a 10 s window for the window to qualify. |
| Freq (Hz) | 4 – 20 | Frequency band of interest. Sets the bandpass filter and the autocorrelation lag-search range. |
| Periodicity ≥ | 0.20 | Minimum normalized autocorrelation peak (0–1) in periodic mode. Ignored in envelope mode. |
| Gap (s) | 30 | Maximum gap between qualifying windows that will still be merged into one event. Bump above ~180 s to merge across consecutive files. |
| Envelope mode | off | When checked, the autocorrelation gates are bypassed. Detection only requires bandpass amplitude and ≥10 peaks per 10 s window. |
| Scan All Channels | — | Run detection. After the first scan, this button stays enabled — tweak parameters and click again to rescan the same target. |
| Scan Baselines… | — | Same as the menu item; opens the folder picker. |
| Full file (top plot) | off | When checked, the raw-signal plot shows the entire TDD file with the detected segment highlighted (instead of zooming on the event). |

## 3.3 Progress and Status

- Analyzing indicator — a red, blinking "● Analyzing…" appears next to the progress bar while a scan is running.

- Progress bar — fills as channels finish; the label shows the latest channel that completed (channel completion order is non-deterministic because of parallelism).

- Status line — announces report writes ("Report written: <name>") and final totals.

## 3.4 Results Table

After scanning, each row is a contiguous segment of step noise on a particular channel. Click a row to plot it; click any column header to sort by that column (an arrow indicator in the header shows current direction).

| Column | Description |
| --- | --- |
| File | Filename of the source TDD (truncated). For cross-file events this is the first file in the merged span. |
| Channel | Hardware channel ID (1-based, sorted numerically). |
| Start (s) | Start of the segment, measured from the beginning of the source file. |
| End (s) | End of the segment. For cross-file events this can exceed the source file's duration — the surplus represents time in subsequent files. |
| Freq (Hz) | Periodic mode: autocorrelation peak frequency. Envelope mode: peaks per second. |
| Amp (mV) | Maximum peak-to-peak of the filtered signal across the merged sub-windows of the segment. |
| Pre BL (mV) | Pre-baseline: maximum raw value in the 10-second window immediately before the segment. |
| Post BL (mV) | Post-baseline: maximum raw value in the 10-second window immediately after the segment. Shows "—" if the segment touches the end of the file. |
| Periodicity | Highest autocorrelation score (0–1) seen across the segment. 0.0 in envelope mode (not measured). |

## 3.5 Plot Panel

- Top — raw signal in millivolts. Detected region highlighted red. Pre/post baseline horizontal dashes shown. Toggle Full file (top plot) to see the whole file.

- Bottom — the same channel through the bandpass filter, focused on the event window so you can see the periodic structure (or burst envelope in envelope mode).

- Both plots include 5 s of context before and after the event by default. The matplotlib toolbar provides zoom, pan, and save.

# 4. Detection Algorithm

## 4.1 Signal Conditioning

Each channel is filtered with a 2nd-order zero-phase Butterworth bandpass. The passband is wider than the user-selected range — 1 Hz below the low cutoff and 5 Hz above the high cutoff — so the rolloff edges don't clip events near the band boundaries. Filter coefficients are designed once per file and shared across all channel threads.

## 4.2 Sliding Window

- Window length: 10 seconds.

- Step: 5 seconds (50 % overlap).

Each window passes through a cascade of gates. A window must clear every gate to qualify.

### Amplitude gate

Peak-to-peak of the filtered window must exceed Amp ≥ (mV).

### Periodic mode (default)

- Autocorrelation periodicity gate — the FFT-based normalized autocorrelation must have a peak ≥ Periodicity ≥ inside the lag range that corresponds to the chosen frequency band (lag = sample_rate / freq).

- Second-harmonic gate — a true periodic signal repeats at multiples of the fundamental lag, so the autocorrelation at 2× the peak lag must be ≥ 0.05. A single transient pulse passes the autocorrelation peak but fails this check.

- Peak-count gate — the count of positive peaks above 20 % of the segment's pp must be at least 30 % of the count expected at the detected frequency. This rejects single spike + ringdown patterns that fool autocorrelation.

- Frequency = sample_rate / peak-lag.

### Envelope mode (toggle)

- Skips the autocorrelation and second-harmonic gates entirely.

- Counts positive peaks (above 20 % of pp) in the window.

- Window passes if there are at least 10 peaks AND the implied frequency (peaks / 10 s) lies inside the user's frequency band.

- Frequency = peaks / 10 s. Periodicity is reported as 0.0 because it is not measured.

## 4.3 Merging and Boundary Refinement

- Adjacent qualifying windows within Gap (s) are merged into one segment. Maximum amplitude and periodicity are kept; frequencies are averaged.

- After merging, the segment edges are refined: the algorithm walks outward from each edge in 2-second sub-windows and extends the segment as long as the local pp is at least 30 % of the segment's measured pp. Up to 3 consecutive quiet sub-windows (≈6 s) are tolerated before stopping, so brief envelope dips don't truncate events.

- Refinement may cause initially separate segments to overlap, so a second merge pass is run afterward.

## 4.4 Cross-file Merging

In a folder scan, after every file in a folder has been processed, the per-file segments are projected onto wall-clock time using each TDD file's DataStartTime. Same-channel segments whose wall-clock gap is within the user's Gap (s) value are merged into a single event. The merged event is reported against the first contributing file, with start/end times measured from that file's start (so end may exceed that file's duration when the event spans into later files).

To merge across files that are recorded ~3 minutes apart, set Gap (s) to roughly the file-to-file gap (e.g. 200) before scanning.

## 4.5 Baselines

- Pre-baseline — the maximum raw value in the 10-second window immediately before the segment.

- Post-baseline — the maximum raw value in the 10-second window immediately after the segment, when data is available; otherwise blank.

# 5. Typical Workflow

## 5.1 Single file

- File → Open TDD… (or Ctrl+O).

- Adjust Amp / Freq / Periodicity / Gap if needed. Try Envelope mode for non-periodic bursts.

- Click Scan All Channels.

- Sort columns to find the strongest events. Click any row to inspect.

## 5.2 One folder

- File → Scan Baselines…

- Browse to the folder with *Baseline*.tdd files. The picker shows the matching files for confirmation.

- Click OK. Scan starts immediately and writes a report to results/ when it finishes.

## 5.3 Many folders

- File → Scan Folders… to add folders one at a time, or File → Scan All Runs… to pick a root and let the app discover every Baseline-containing folder under it.

- Inspect the discovered runs in the preview dialog. Remove any you don't want.

- Scan. One report is written per source folder, named after the run and the time of the scan.

## 5.4 Re-scanning with new parameters

After any scan, the Scan All Channels button stays enabled. Tweak the toolbar parameters and click again — the same files are rescanned with the new settings. There is no need to re-pick the folder.

# 6. Output Files

Reports are written under the project's results/ directory (created if it does not exist). One pair of files is produced per source folder:

- <run-name>_<timestamp>.txt — a human-readable text report listing every detected segment with its file, channel, start/end, frequency, amplitude, baselines, and periodicity. The header records the parameters used for the scan.

- <run-name>_<timestamp>.json — a Sherlock-compatible RunResults JSON, suitable for downstream analysis. Run-level metadata (run_id, wafer_id, die_number, lot_number, instrument_id) is pulled from the TDD header. Aggressors are keyed by channel, then by wall-clock ISO timestamp of the event's first sample.

The run name is derived from the first four dash-separated parts of the TDD filename (e.g. HAK04-008B-02L64869w15-205B16a). The File column inside the .txt report shows only the 14-digit datetime code from each filename.

## 6.1 Settings file

Toolbar parameters are stored in step_noise_finder_config.json next to the app. It is rewritten every time you change a value. To reset to defaults, delete the file.

# 7. Parameter Tuning

| Parameter | Increase to… | Decrease to… | Note |
| --- | --- | --- | --- |
| Amp ≥ (mV) | Reduce false positives from low-amplitude noise | Catch weaker step noise | 100 mV is typical for Nabsys instruments; try 50 mV if events look clean but aren't reported. |
| Freq range | Widen to catch variable-frequency steps | Narrow to target a known frequency | 4–20 Hz covers most observed step noise. |
| Periodicity | Demand cleaner periodic signals (fewer false positives) | Catch irregular or amplitude-modulated events | 0.20 default is fairly permissive; bump to 0.4+ for strict. |
| Gap (s) | Merge events across long quiet stretches or across files | Keep events more granular | Set to ~200 to merge across 30-min files recorded ≈3 min apart. |
| Envelope mode | (checkbox — enable for bursty / non-periodic events) | (uncheck for strict periodic detection) | Best for ragged spike trains in the band; reports peaks/sec as the frequency. |

# 8. Technical Details

## 8.1 TDD format

- Magic 0x5342414E ("NABS", little-endian).

- Per-channel int16 samples interleaved across time.

- Per-channel calibration factors applied during load to convert raw counts to microvolts.

- Header fields used by reports include RunID, DetectorWaferID, DetectorDieNumber, DetectorLotNumber, SystemID, and DataStartTime.

## 8.2 Channel ordering

Channels are recorded in hardware-scrambled order. The reader sorts them numerically (1–2…256) for display and analysis.

## 8.3 Performance

- FFT autocorrelation — scipy.fft.rfft/irfft replaces direct convolution; ~1700× faster on the inner autocorrelation loop than np.correlate, with float-epsilon-equivalent results.

- Channel-level parallelism — ThreadPoolExecutor with min(cpu_count, 8) workers. NumPy/SciPy release the GIL, so Python threads give real parallelism here.

- LRU file cache — bounded by total waveform bytes (8 GB). Older files are evicted automatically as new ones load. If the user clicks a row whose file was evicted, the file is reloaded transparently to render the plot.

## 8.4 Threading and Tk safety

All scanning runs in background threads. Progress and table updates dispatch back to the Tk main thread via root.after(0, …), the only thread-safe Tk update path.

## 8.5 Dependencies

- Python 3.13 (the installed venv version).

- numpy, scipy, matplotlib.

- Optional: sherlock (Nabsys-Organization/sherlock). When installed, RunResults JSON output is enabled.

## 8.6 Source layout

- src/step_finder.py — main GUI and detection.

- src/tdd_reader.py — TDD v3.1 binary parser.

- src/plot_results.py — helper to plot channel-vs-step-noise across all results/*.json.

- doc/gen_finder_guide.js — source for this user guide; run with node to regenerate the docx.