File Description: [SampleID]-[System ID]-[Module Number]-[YYYYMMDDHHMMSS].tdd

File format version: 3

## Introduction

File containing interlaced time domain data from X number of channels. File format now denote a major and minor revision. The first number will match the “File Version”. The second number will track revisions as this document is edited towards a new version.

File created by:

- System software

### File used by:

- Signal processing software

## Data Type Notes

- All header data of “var” length are prepended with a 4 byte integer which indicates the length of the field following.

- Tstmp is a timestamp defined as the number of seconds that have elapsed since 01/01/1904 00:00:00.00 UTC, using the Gregorian calendar and ignoring leap seconds. It is stored as two 64 bit integers. The first is an unsigned 64 bit integer which represents the fractional portion, multiplied by 264. The second is a signed 64 bit integer representing whole seconds. The full timestamp is a sum of the 2 numbers.

- All Int’s are signed, unless otherwise noted.

## File Format

Binary data is stored in little-endian order.

File contains a header section followed by interlaced channel data.

### Header section:

| Description | Type | Len (bytes) | Notes |
| --- | --- | --- | --- |
| Nabsys file | Int | 4 | 0x5342414E = (‘N’ ‘A’ ‘B’ ‘S’) |
| File type | Int | 4 | 0x00011170 = 70000 |
| File version | Int | 4 | 3 |
| Channel Count | Int | 4 | Number of channels of data in this file. Important when decimating data. |
| File Number | Int | 4 | Since runs can produce more than one file, this number represents the order this file was produced (starting with 1) |
| Module Number | Int | 4 | Module position on the system |
| Run Start Time | Tstmp | 16 | Start time of the run (not the start of the data collection) |
| System ID | Str | var | System ID as defined in the system software |
| Software | Str | var | System software name |
| Software Version | Str | var | System software version [Major].[Minor].[Fix].[Build] |
| System Controller SW Version | Str | var | System Controller (SOM) software version [Major].[Minor].[Fix].[Build] |
| Operator | Str | var | Operator account name |
| Run ID | Str | var | YYYYMMDD_HHMMSS-[System ID]-[Module Number] |
| Protocol Name | Str | var | Name of the protocol used to generate the data |
| Protocol Version | Str | var | Version of the protocol that was run |
| Settings Group | Str | var | Name of the group of settings used for the protocol |
| Settings Version | Str | var | Version of the settings file that was used |
| Module Model Number | Str | var | Model number of the module used to collect this data. |
| Module Serial Number | Str | var | Serial number assigned to the module the data was collected on |
| Detector Type | Str | var | Type identifier for the detector |
| Detector Lot Number | Str | var | Typically a 7 digit number |
| Detector Wafer ID | Str | Var | Typically a 1 digit number |
| Detector Die Number | Str | var | Typically a letter representing the row, followed by a 2 digit number representing the column. |
| Reagent Lot | Str | var |  |
| Sample ID | Str | var | Customer defined sample ID |
| Scientist Name | Str | var | Place to store the name of the scientist responsible for the sample |
| Sample Type | Str | var | Field to indicate which instrument sample type definition was used |
| Sample Description | Str | var |  |
| Sample Method | Str | var |  |
| Tag Size | Str | var |  |
| Sample Rate | Int | 4 | Collection rate. Sample rate is identical for each channel. |
| Factor | Array[Dbl] | 8 x #chan | Double precision floating point number to multiply by the data to receive data points in microvolts. Stored as an array with an element for each channel of the detector. |
| Low Pass Filter | Array[Dbl] | 8 x #chan | Low pass cutoff frequency. Stored as an array with an element for each channel of the detector. |
| High Pass Filter | Array[Dbl] | 8 x #chan | High pass cutoff frequency. Stored as an array with an element for each channel of the detector. |
| Channel IDs | Array[Int] | 4 x #chan | An ID is assigned for each channel in the system. This determines which channel(s) the data was collected on. |
| Segment Length | Int | 8 | If the data is a collection of segments, this defines the segment length in sample counts. Each channel must be collected in identical segment lengths. |
| Data Start Time | Tstmp | 16 | Time zero of the data |

Table 1 Header Structure

### Data:

- Interleaved data of all channels

| Description | Data Type | Len (bytes) | Notes |
| --- | --- | --- | --- |
| Channel data | Int | 2 | Magnitude of the channel differential. Multiply by “Factor” to convert to microvolts. |

Table 2 - Data

## Documentation File Revision History

Rev 1, 13Jun2017, T. Hutchins, M. Kaiser

- File created

- Approved, T.Hutchins

Rev1.1, 01May2018, T.Hutchins, M.Kaiser

- Eliminated the “Flour Label” meta data field

Rev2.0, 26Oct2018, T.Hutchins

- Removed the extraneous reagent lot fields

- Added a field for system controller software version

Rev3.0, 01Nov2019, T.Hutchins, M.Kaiser

- Added meta data for sample type and protocol/settings versions

Rev3.1, 15Jun2020, M.Kaiser

- Documentation corrections

- Changed file version from “2” to “3”

- Changed footer from “file format: 2” to “file format: 3”